## SST
Q-Exactive System Suitability Test (SST): Overview sheet
Project: | Surgery |  |  |  | Resol. ≥ 60'000 |  | 79'800 / 79'700 | at m/z 201 (S22, 4.35 min, neg) / m/z 205 (T32, 2.7 min, pos)
Matrix: | Jejunum biopsies whole wall & mucosa, Serum |  |  |  | SD RT ≤ 0.05 min |  | ok
LC-MS sequence: | 20230203_SurgeryTest_Biops_Serum_FAO_Exp_LC |  |  |  | Dm/z ≤ 5 ppm pos / ≤ 9 ppm neg |  | ok
Column No.: | 121 |  |  |  | Intensity ≥ Threshold (E+06) |  | ok
Date/Initials: | 07.02.23 / KF |  |  |  | CV Intensity ≤ 15% |  | (ok)
[5 empty] | SST passed |  | yes
| Name | Formula | Replicate Name | RT (min) | SD RT (min) | Mass Error (ppm) | Peak Area | CV Area | Peak Height | CV Height
Negative ionization [M-H] | A44_Acetoacetic acid | C4H6O3 | Neg_SST_1 | 1.16 |  | 0.7 | 27666934 |  | 10896337
| A44_Acetoacetic acid | C4H6O3 | Neg_SST_2 | 1.16 |  | 0.6 | 27592858 |  | 10847648
| A44_Acetoacetic acid | C4H6O3 | Neg_SST_3 | 1.17 | 0.0057735 | 0.7 | 30735416 | 6.256% | 12168396 | 6.625%
| A44_Acetoacetic acid | C4H6O3 | Neg_SST_4 | 1.18 | 0.0095743 | 0.6 | 40291560 | 18.99% | 15629173 | 18.14%
| D33_Decanoic acid | C10H20O2 | Neg_SST_1 | 5.74 |  | -0.1 | 5023572 |  | 2004188
| D33_Decanoic acid | C10H20O2 | Neg_SST_2 | 5.75 |  | 0 | 4977404 |  | 1932836
| D33_Decanoic acid | C10H20O2 | Neg_SST_3 | 5.74 | 0.0057735 | 0.2 | 6011615 | 10.95% | 2747907 | 20.26%
| D33_Decanoic acid | C10H20O2 | Neg_SST_4 | 5.74 | 0.005 | 0 | 6255475 | 11.89% | 2346398 | 16.53%
| G44_L-Glutamic acid | C5H9NO4 | Neg_SST_1 | 0.72 |  | 0.4 | 25284364 |  | 10910981
| G44_L-Glutamic acid | C5H9NO4 | Neg_SST_2 | 0.72 |  | 0.4 | 26852324 |  | 11385136
| G44_L-Glutamic acid | C5H9NO4 | Neg_SST_3 | 0.72 | 1.3597e-16 | 0.5 | 27307696 | 4.008% | 11231270 | 2.164%
| G44_L-Glutamic acid | C5H9NO4 | Neg_SST_4 | 0.72 | 0 | 0.9 | 35506556 | 15.99% | 13503213 | 10.04%
| K03_Kynurenic acid | C10H7NO3 | Neg_SST_1 | 3.02 |  | -3.5 | 15150266 |  | 5626714
| K03_Kynurenic acid | C10H7NO3 | Neg_SST_2 | 3.02 |  | -3.3 | 16303660 |  | 5954423
| K03_Kynurenic acid | C10H7NO3 | Neg_SST_3 | 3.03 | 0.0057735 | -3.3 | 17940488 | 8.516% | 6216604 | 4.982%
| K03_Kynurenic acid | C10H7NO3 | Neg_SST_4 | 3.05 | 0.014142 | -3.8 | 25807152 | 25.58% | 9905443 | 28.89%
| S22_Sebacic acid | C10H18O4 | Neg_SST_1 | 4.35 |  | -4.8 | 191907328 |  | 79740968
| S22_Sebacic acid | C10H18O4 | Neg_SST_2 | 4.35 |  | -4.6 | 217564800 |  | 92004408
| S22_Sebacic acid | C10H18O4 | Neg_SST_3 | 4.35 | 0 | -4.9 | 241948368 | 11.52% | 101063448 | 11.77%
| S22_Sebacic acid | C10H18O4 | Neg_SST_4 | 4.35 | 0 | -4.9 | 282282080 | 16.47% | 114637416 | 15.2%
| T32_L-Tryptophan | C11H12N2O2 | Neg_SST_1 | 2.89 |  | -5.3 | 115183720 |  | 48694632
| T32_L-Tryptophan | C11H12N2O2 | Neg_SST_2 | 2.87 |  | -4.7 | 101918856 |  | 43408360
| T32_L-Tryptophan | C11H12N2O2 | Neg_SST_3 | 2.9 | 0.015275 | -4.7 | 104076896 | 6.648% | 40818132 | 9.06%
| T32_L-Tryptophan | C11H12N2O2 | Neg_SST_4 | 2.87 | 0.015 | -5.2 | 164402384 | 24.1% | 51584064 | 10.62%
| T34_L-Tyrosine | C9H11NO3 | Neg_SST_1 | 1.03 |  | -2 | 22597836 |  | 12102063
| T34_L-Tyrosine | C9H11NO3 | Neg_SST_2 | 1.03 |  | -2.1 | 21738170 |  | 10513619
| T34_L-Tyrosine | C9H11NO3 | Neg_SST_3 | 1.03 | 0 | -2.1 | 25107732 | 7.564% | 12166232 | 8.075%
| T34_L-Tyrosine | C9H11NO3 | Neg_SST_4 | 1.02 | 0.005 | -2.1 | 29423680 | 13.95% | 15070075 | 15.23%
Positive ionization [M+H] | G44_L-Glutamic acid | C5H9NO4 | Pos_SST_1 | 0.75 |  | -0.6 | 225643824 |  | 88179264
| G44_L-Glutamic acid | C5H9NO4 | Pos_SST_2 | 0.75 |  | -0.5 | 230621120 |  | 91395712
| G44_L-Glutamic acid | C5H9NO4 | Pos_SST_3 | 0.75 | 0 | -0.4 | 230719584 | 1.267% | 90832848 | 1.906%
| G44_L-Glutamic acid | C5H9NO4 | Pos_SST_4 | 0.77 | 0.01 | -0.3 | 209286128 | 4.523% | 63429344 | 16.09%
| K03_Kynurenic acid | C10H7NO3 | Pos_SST_1 | 3.02 |  | -0.3 | 208679856 |  | 74251344
| K03_Kynurenic acid | C10H7NO3 | Pos_SST_2 | 3.05 |  | -0.2 | 210620336 |  | 77991232
| K03_Kynurenic acid | C10H7NO3 | Pos_SST_3 | 3.02 | 0.017321 | 0 | 210391040 | 0.5052% | 74536176 | 2.754%
| K03_Kynurenic acid | C10H7NO3 | Pos_SST_4 | 3.04 | 0.015 | 0.1 | 221655120 | 2.792% | 80700392 | 3.991%
| K04_L-Kynurenine | C10H12N2O3 | Pos_SST_1 | 1.95 |  | -0.7 | 457546464 |  | 221890608
| K04_L-Kynurenine | C10H12N2O3 | Pos_SST_2 | 1.96 |  | -0.5 | 459831456 |  | 218511840
| K04_L-Kynurenine | C10H12N2O3 | Pos_SST_3 | 1.96 | 0.0057735 | -0.3 | 454089216 | 0.6324% | 219068784 | 0.8241%
| K04_L-Kynurenine | C10H12N2O3 | Pos_SST_4 | 1.95 | 0.0057735 | -0.1 | 473024480 | 1.795% | 226347952 | 1.617%
| T32_L-Tryptophan | C11H12N2O2 | Pos_SST_1 | 2.84 |  | -0.6 | 599210752 |  | 189970576
| T32_L-Tryptophan | C11H12N2O2 | Pos_SST_2 | 2.86 |  | -0.5 | 622308928 |  | 165750400
| T32_L-Tryptophan | C11H12N2O2 | Pos_SST_3 | 2.84 | 0.011547 | -0.5 | 593711232 | 2.508% | 183279696 | 6.962%
| T32_L-Tryptophan | C11H12N2O2 | Pos_SST_4 | 2.82 | 0.01633 | -0.3 | 645095296 | 3.826% | 158891904 | 8.349%
| T34_L-Tyrosine | C9H11NO3 | Pos_SST_1 | 1.02 |  | -0.5 | 105686480 |  | 54008368
| T34_L-Tyrosine | C9H11NO3 | Pos_SST_2 | 1.03 |  | -0.1 | 102129904 |  | 53065176
| T34_L-Tyrosine | C9H11NO3 | Pos_SST_3 | 1.03 | 0.0057735 | -0.1 | 103136968 | 1.769% | 53375916 | 0.8987%
| T34_L-Tyrosine | C9H11NO3 | Pos_SST_4 | 1.02 | 0.0057735 | 0 | 112160112 | 4.264% | 56249232 | 2.654%
PASTE SKYLINE REPORT IN YELLOW CELL BELOW
| Molecule Name | Molecule Formula | Precursor Adduct | Precursor Mz | Replicate Name | Best Retention Time | Mean Best Retention Time | Stdev Best Retention Time | Average Mass Error PPM | Total Area MS1 | Mean Total Area | Cv Total Area | Max Height | Mean Max Height | Cv Max Height
| A44_Acetoacetic acid | C4H6O3 | [M-H] | 101.0244 | Neg_SST_1 | 1.16 | 1.17 | 0.01 | 0.7 | 27666934 | 31571692 | 19% | 10896337 | 12385389 | 18.1%
| A44_Acetoacetic acid | C4H6O3 | [M-H] | 101.0244 | Neg_SST_2 | 1.16 | 1.17 | 0.01 | 0.6 | 27592858 | 31571692 | 19% | 10847648 | 12385389 | 18.1%
| A44_Acetoacetic acid | C4H6O3 | [M-H] | 101.0244 | Neg_SST_3 | 1.17 | 1.17 | 0.01 | 0.7 | 30735416 | 31571692 | 19% | 12168396 | 12385389 | 18.1%
| A44_Acetoacetic acid | C4H6O3 | [M-H] | 101.0244 | Neg_SST_4 | 1.18 | 1.17 | 0.01 | 0.6 | 40291560 | 31571692 | 19% | 15629173 | 12385389 | 18.1%
| D33_Decanoic acid | C10H20O2 | [M-H] | 171.1391 | Neg_SST_1 | 5.74 | 5.75 | 0 | -0.1 | 5023572 | 5567016 | 11.9% | 2004188 | 2257832 | 16.5%
| D33_Decanoic acid | C10H20O2 | [M-H] | 171.1391 | Neg_SST_2 | 5.75 | 5.75 | 0 | 0 | 4977404 | 5567016 | 11.9% | 1932836 | 2257832 | 16.5%
| D33_Decanoic acid | C10H20O2 | [M-H] | 171.1391 | Neg_SST_3 | 5.74 | 5.75 | 0 | 0.2 | 6011615 | 5567016 | 11.9% | 2747907 | 2257832 | 16.5%
| D33_Decanoic acid | C10H20O2 | [M-H] | 171.1391 | Neg_SST_4 | 5.74 | 5.75 | 0 | 0 | 6255475 | 5567016 | 11.9% | 2346398 | 2257832 | 16.5%
| G44_L-Glutamic acid | C5H9NO4 | [M-H] | 146.0459 | Neg_SST_1 | 0.72 | 0.72 | 0 | 0.4 | 25284364 | 28737735 | 16% | 10910981 | 11757650 | 10%
| G44_L-Glutamic acid | C5H9NO4 | [M-H] | 146.0459 | Neg_SST_2 | 0.72 | 0.72 | 0 | 0.4 | 26852324 | 28737735 | 16% | 11385136 | 11757650 | 10%
| G44_L-Glutamic acid | C5H9NO4 | [M-H] | 146.0459 | Neg_SST_3 | 0.72 | 0.72 | 0 | 0.5 | 27307696 | 28737735 | 16% | 11231270 | 11757650 | 10%
| G44_L-Glutamic acid | C5H9NO4 | [M-H] | 146.0459 | Neg_SST_4 | 0.72 | 0.72 | 0 | 0.9 | 35506556 | 28737735 | 16% | 13503213 | 11757650 | 10%
| K03_Kynurenic acid | C10H7NO3 | [M-H] | 188.0353 | Neg_SST_1 | 3.02 | 3.03 | 0.01 | -3.5 | 15150266 | 18800392 | 25.6% | 5626714 | 6925796 | 28.9%
| K03_Kynurenic acid | C10H7NO3 | [M-H] | 188.0353 | Neg_SST_2 | 3.02 | 3.03 | 0.01 | -3.3 | 16303660 | 18800392 | 25.6% | 5954423 | 6925796 | 28.9%
| K03_Kynurenic acid | C10H7NO3 | [M-H] | 188.0353 | Neg_SST_3 | 3.03 | 3.03 | 0.01 | -3.3 | 17940488 | 18800392 | 25.6% | 6216604 | 6925796 | 28.9%
| K03_Kynurenic acid | C10H7NO3 | [M-H] | 188.0353 | Neg_SST_4 | 3.05 | 3.03 | 0.01 | -3.8 | 25807152 | 18800392 | 25.6% | 9905443 | 6925796 | 28.9%
| S22_Sebacic acid | C10H18O4 | [M-H] | 201.1132 | Neg_SST_1 | 4.35 | 4.35 | 0 | -4.8 | 191907328 | 233425644 | 16.5% | 79740968 | 96861560 | 15.2%
| S22_Sebacic acid | C10H18O4 | [M-H] | 201.1132 | Neg_SST_2 | 4.35 | 4.35 | 0 | -4.6 | 217564800 | 233425644 | 16.5% | 92004408 | 96861560 | 15.2%
| S22_Sebacic acid | C10H18O4 | [M-H] | 201.1132 | Neg_SST_3 | 4.35 | 4.35 | 0 | -4.9 | 241948368 | 233425644 | 16.5% | 101063448 | 96861560 | 15.2%
| S22_Sebacic acid | C10H18O4 | [M-H] | 201.1132 | Neg_SST_4 | 4.35 | 4.35 | 0 | -4.9 | 282282080 | 233425644 | 16.5% | 114637416 | 96861560 | 15.2%
| T32_L-Tryptophan | C11H12N2O2 | [M-H] | 203.0826 | Neg_SST_1 | 2.89 | 2.88 | 0.01 | -5.3 | 115183720 | 121395464 | 24.1% | 48694632 | 46126297 | 10.6%
| T32_L-Tryptophan | C11H12N2O2 | [M-H] | 203.0826 | Neg_SST_2 | 2.87 | 2.88 | 0.01 | -4.7 | 101918856 | 121395464 | 24.1% | 43408360 | 46126297 | 10.6%
| T32_L-Tryptophan | C11H12N2O2 | [M-H] | 203.0826 | Neg_SST_3 | 2.9 | 2.88 | 0.01 | -4.7 | 104076896 | 121395464 | 24.1% | 40818132 | 46126297 | 10.6%
| T32_L-Tryptophan | C11H12N2O2 | [M-H] | 203.0826 | Neg_SST_4 | 2.87 | 2.88 | 0.01 | -5.2 | 164402384 | 121395464 | 24.1% | 51584064 | 46126297 | 10.6%
| T34_L-Tyrosine | C9H11NO3 | [M-H] | 180.0666 | Neg_SST_1 | 1.03 | 1.03 | 0 | -2 | 22597836 | 24716855 | 14% | 12102063 | 12462997 | 15.2%
| T34_L-Tyrosine | C9H11NO3 | [M-H] | 180.0666 | Neg_SST_2 | 1.03 | 1.03 | 0 | -2.1 | 21738170 | 24716855 | 14% | 10513619 | 12462997 | 15.2%
| T34_L-Tyrosine | C9H11NO3 | [M-H] | 180.0666 | Neg_SST_3 | 1.03 | 1.03 | 0 | -2.1 | 25107732 | 24716855 | 14% | 12166232 | 12462997 | 15.2%
| T34_L-Tyrosine | C9H11NO3 | [M-H] | 180.0666 | Neg_SST_4 | 1.02 | 1.03 | 0 | -2.1 | 29423680 | 24716855 | 14% | 15070075 | 12462997 | 15.2%
| G44_L-Glutamic acid | C5H9NO4 | [M+H] | 148.0604 | Pos_SST_1 | 0.75 | 0.76 | 0.01 | -0.6 | 225643824 | 224067664 | 4.5% | 88179264 | 83459292 | 16.1%
| G44_L-Glutamic acid | C5H9NO4 | [M+H] | 148.0604 | Pos_SST_2 | 0.75 | 0.76 | 0.01 | -0.5 | 230621120 | 224067664 | 4.5% | 91395712 | 83459292 | 16.1%
| G44_L-Glutamic acid | C5H9NO4 | [M+H] | 148.0604 | Pos_SST_3 | 0.75 | 0.76 | 0.01 | -0.4 | 230719584 | 224067664 | 4.5% | 90832848 | 83459292 | 16.1%
| G44_L-Glutamic acid | C5H9NO4 | [M+H] | 148.0604 | Pos_SST_4 | 0.77 | 0.76 | 0.01 | -0.3 | 209286128 | 224067664 | 4.5% | 63429344 | 83459292 | 16.1%
| K03_Kynurenic acid | C10H7NO3 | [M+H] | 190.0499 | Pos_SST_1 | 3.02 | 3.04 | 0.02 | -0.3 | 208679856 | 212836588 | 2.8% | 74251344 | 76869786 | 4%
| K03_Kynurenic acid | C10H7NO3 | [M+H] | 190.0499 | Pos_SST_2 | 3.05 | 3.04 | 0.02 | -0.2 | 210620336 | 212836588 | 2.8% | 77991232 | 76869786 | 4%
| K03_Kynurenic acid | C10H7NO3 | [M+H] | 190.0499 | Pos_SST_3 | 3.02 | 3.04 | 0.02 | 0 | 210391040 | 212836588 | 2.8% | 74536176 | 76869786 | 4%
| K03_Kynurenic acid | C10H7NO3 | [M+H] | 190.0499 | Pos_SST_4 | 3.04 | 3.04 | 0.02 | 0.1 | 221655120 | 212836588 | 2.8% | 80700392 | 76869786 | 4%
| K04_L-Kynurenine | C10H12N2O3 | [M+H] | 209.0921 | Pos_SST_1 | 1.95 | 1.95 | 0 | -0.7 | 457546464 | 461122904 | 1.8% | 221890608 | 221454796 | 1.6%
| K04_L-Kynurenine | C10H12N2O3 | [M+H] | 209.0921 | Pos_SST_2 | 1.96 | 1.95 | 0 | -0.5 | 459831456 | 461122904 | 1.8% | 218511840 | 221454796 | 1.6%
| K04_L-Kynurenine | C10H12N2O3 | [M+H] | 209.0921 | Pos_SST_3 | 1.96 | 1.95 | 0 | -0.3 | 454089216 | 461122904 | 1.8% | 219068784 | 221454796 | 1.6%
| K04_L-Kynurenine | C10H12N2O3 | [M+H] | 209.0921 | Pos_SST_4 | 1.95 | 1.95 | 0 | -0.1 | 473024480 | 461122904 | 1.8% | 226347952 | 221454796 | 1.6%
| T32_L-Tryptophan | C11H12N2O2 | [M+H] | 205.0972 | Pos_SST_1 | 2.84 | 2.84 | 0.02 | -0.6 | 599210752 | 615081552 | 3.8% | 189970576 | 174473144 | 8.3%
| T32_L-Tryptophan | C11H12N2O2 | [M+H] | 205.0972 | Pos_SST_2 | 2.86 | 2.84 | 0.02 | -0.5 | 622308928 | 615081552 | 3.8% | 165750400 | 174473144 | 8.3%
| T32_L-Tryptophan | C11H12N2O2 | [M+H] | 205.0972 | Pos_SST_3 | 2.84 | 2.84 | 0.02 | -0.5 | 593711232 | 615081552 | 3.8% | 183279696 | 174473144 | 8.3%
| T32_L-Tryptophan | C11H12N2O2 | [M+H] | 205.0972 | Pos_SST_4 | 2.82 | 2.84 | 0.02 | -0.3 | 645095296 | 615081552 | 3.8% | 158891904 | 174473144 | 8.3%
| T34_L-Tyrosine | C9H11NO3 | [M+H] | 182.0812 | Pos_SST_1 | 1.02 | 1.02 | 0 | -0.5 | 105686480 | 105778366 | 4.3% | 54008368 | 54174673 | 2.7%
| T34_L-Tyrosine | C9H11NO3 | [M+H] | 182.0812 | Pos_SST_2 | 1.03 | 1.02 | 0 | -0.1 | 102129904 | 105778366 | 4.3% | 53065176 | 54174673 | 2.7%
| T34_L-Tyrosine | C9H11NO3 | [M+H] | 182.0812 | Pos_SST_3 | 1.03 | 1.02 | 0 | -0.1 | 103136968 | 105778366 | 4.3% | 53375916 | 54174673 | 2.7%
| T34_L-Tyrosine | C9H11NO3 | [M+H] | 182.0812 | Pos_SST_4 | 1.02 | 1.02 | 0 | 0 | 112160112 | 105778366 | 4.3% | 56249232 | 54174673 | 2.7%

## Chart
Q-Exactive System Suitability Test (SST): Values to be reported
Report the average RT and intensity measured for each molecular ion during your LC-MS analysis in the progress chart
[5 empty] | A44_Acetoacetic acid |  |  | D33_Decanoic acid |  |  | G44_L-Glutamic acid |  |  | K03_Kynurenic acid |  |  | S22_Sebacic acid |  |  | T32_L-Tryptophan |  |  | T34_L-Tyrosine |  |  | G44_L-Glutamic acid |  |  | K03_Kynurenic acid |  |  | K04_L-Kynurenine |  |  | T32_L-Tryptophan |  |  | T34_L-Tyrosine
Comments | SST from | Column | Date | Resolution (-/+) | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area | RT (min) | Height | Area
Surgery test 1 / biopsies & serum | 14.12.22 | 121 | 07.02.23 / KF | 79'800 / 79'700 | 1.17 | 12385389 | 31571692 | 5.75 | 2257832 | 5567016 | 0.72 | 11757650 | 28737735 | 3.03 | 6925796 | 18800392 | 4.35 | 96861560 | 233425644 | 2.88 | 46126297 | 121395464 | 1.03 | 12462997 | 24716855 | 0.76 | 83459292 | 224067664 | 3.04 | 76869786 | 212836588 | 1.95 | 221454796 | 461122904 | 2.84 | 174473144 | 615081552 | 1.02 | 54174673 | 105778366