|  |
| --- |
| STANDARD OPERATING PROCEDURE |
|

|  |
| --- |
| **Title: Liquid Chromatography, Waters nanoACQUITY UPLC system**  |
|  |  |
| **Version #: 2** | **Author: PNNL lab** |
| **Date: 09/01/2015** |  |

 |

# Purpose

The purpose of this document is to describe the Liquid Chromatography (LC) method for development of CPTAC MRM assays.

# Scope

This procedure is designed to help the setup of LC gradient and method parameters on Waters nanoACQUITY UPLC system.

# Responsibilities

It is the responsibility of person(s) performing this procedure to be familiar with laboratory safety procedures. The interpretation of results must be done by a person trained in the procedure and familiar with such interpretation.

# Equipment

UPLC: nanoACQUITY (Waters, Part Number 176016000)

# Materials

Injection loop: 5 uL peeksil

Column: 100 μm x 100 mm, BEH C18, 1.7 μm 130 Å (Waters, Part Number 186007485)

Water, HPLC grade (H2O)

Acetonitrile, HPLC grade (ACN) (Fisher Scientific, A955-4)

Formic Acid (0.1%)/Acetonitrile (EMD, FX0437P-1)

Formic Acid (FA) (Agilent Technologies, G2453-85060)

# Reagents

Mobile phase A: 0.1% FA in H2O

Mobile phase B: 0.1% FA in ACN

Strong Needle Wash buffer: 100% ACN

Weak Needle Wash buffer: 0.1% FA in H2O

Seal Wash buffer: 10% ACN in H2O

# Procedure

**Instrument Configuration**

1. **Autosampler parameters:**

**Under General tab:**

‘Partial loop’ under ‘sample loop option’

Wash Solvent: Weak 600 ul, Strong 200 ul

Column Temperature: 42 ºC

Autosampler temperature: 4 ºC

Run time: 8 min

**Under Event tab:**

 Check ‘run events’ box:

|  |  |  |
| --- | --- | --- |
| Time (min) | Event | Action |
| 10.00 | Switch 1 | On |
| 10.01 | Switch 1 | Off |

1. LC gradient methods:

|  |  |  |  |
| --- | --- | --- | --- |
| Time (min) | %Mobile Phase A | %Mobile phase B | Flow rate (ul/min) |
| 0 | 99.5 | 0.5 | 0.5 |
| 9 | 95 | 5 | 0.5 |
| 13 | 92 | 8 | 0.5 |
| 20 | 90 | 10 | 0.5 |
| 26 | 87.5 | 12.5 | 0.5 |
| 34 | 85.5 | 14.5 | 0.5 |
| 48 | 78 | 22 | 0.5 |
| 52 | 65 | 35 | 0.5 |
| 54 | 5 | 95 | 0.5 |
| 60 | 5 | 95 | 0.75 |
| 62 | 90 | 10 | 0.75 |
| 64 | 5 | 95 | 0.75 |
| 65 | 5 | 95 | 0.75 |
| 67 | 99.5 | 0.5 | 0.75 |
| 75 | 99.5 | 0.5 | 0.5 |

# Referenced Documents

List any publications or documents referenced in the SOP.